Otolith Laser Ablation Analysis was performed over several years. All otoliths for a given river/year combination were analyzed together over the course of multiple months. For the most part, datasets were analyzed in corresponding temporal order; meaning sample sets were processed in the order of the year they were collected, generally a year or so after the collection occurred.

87Sr/86Sr in otolith sections mounted on sildes were measured along a trasect from the core to the marine outer section using a laser ablation system (Excite, Teledyne Photon Machines) using a a laser beam size of 65 mm, a fluence of 2.82 J.cm-2, and a scanning speed of 2 mm.sec-1 . 87Sr/86Sr ratios were corrected for mass bias using an exponential law, normalizing to 86Sr/86Sr = 0.1194 (Steiger and Jäger 1977). Isobaric interferences on the 87Sr/86Sr ratios, from 87Rb and 86Kr, were corrected by simultaneously monitoring 85Rb and 83Kr using the corresponding invariant ratios of 87Rb/85Rb = 0.385706 and 86Kr/83Kr = 1.502522 (Steiger and Jäger 1977). Standard reference material SRM 987 (SrCO3, National Institute of Standards and Technology, Gaithersburg, MD, US) was used to verify the quality of data. On one hand, SRM 987 aliquots in 2 M HNO3, or “column standards”, were purified and run with the samples. In addition, a solution of SRM 987 in 5% HNO3 was run repeatedly with samples (1 standard every 3 samples). 87Sr/86Sr long term average and precision (SD 1s) for SRM 987 in our lab is 0.710295 ± 0.000016. Every 87Sr/86Sr values obtained for column standards was within 2 SD from our long-term average. Analytical precision (standard error) of 87Sr/86Sr for all samples was <0.00001. The instruments are in filtered air positive pressure labs, and sample handling and dilutions were performed in laminar flow benches using calibrated balances and pipettors. Aliquots of 2 M HNO3 (Trace Metal Grade), or “column blanks”, were purified and run with samples and standards, producing intensities on average 500 times lower.

Steiger, R. H., Jager, E., 1977. Subcommission on geochronology: Convention on the use of decay constants in geo- and cosmochronology. Earth Planetary Science Letters, 36, 359-362.